## Slide 1: Understanding the Requirements of ISO/IEC 17025:2017

## Slide 2: Introduction
ISO/IEC 17025:2017 is the latest standard of the requirements for the competence of testing and calibration laboratories
***
Previous versions included:
ISO/IEC 17025:2005
ISO/IEC 17025:1999
Notes: ISO/IEC 17025:2017 is the latest standard of the requirements for the competence of testing and calibration laboratories
ISO/IEC 17025 was initially published in 1999. A revision was added in 2005 and the standard was recently updated in November 2017.

## Slide 3: Topics Covered
1. Basics of a Lab-MS and ISO/IEC 17025
What is an Lab-MS?
Benefits of Accreditation
What is needed for the Accreditation Process?
Clause 5 – Structural Requirements
Clause 6 - Resource Requirements
Clause 7 – Process Requirements
Clause 8 – Management System Requirements
2. ISO/IEC 17025:2017 Requirements
Clause 4 – General Requirements
What is ISO/IEC 17025?
3
Notes: Today we will cover the following topics so that you will better understand your company’s Laboratory Management System.
What is ISO/IEC 17025 and what is a Lab-MS
What are the benefits of achieving accreditation
What are the elements necessary to establish and manage a ISO/IEC 17025 Lab-MS, and
What is needed for the ISO/IEC 17025 accreditation process
And finally, we will go through the requirements in the clauses of ISO/IEC 17025:2017

## Slide 4: Section 1: Basics of a Lab-MS What is ISO/IEC 17025:2017 ?
What is ISO/IEC 17025?
What is a Lab-MS?
Benefits of ISO/IEC 17025 accreditation
What is required for the accreditation process?
4
Notes: Now we will explain a Lab-MS and ISO/IEC 17025:2017 in more detail…

## Slide 5: What is ISO/IEC 17025:2017
Outlines the basic elements of a laboratory management system (Lab-MS) which support the defined testing or calibration scope of accreditation
Applies to any organization throughout the world performing any testing or calibration
Does not mandate across-the-board criteria a company must meet, like a certain “level of quality”
Does not “rate” your company against others – but proficiency testing reassures both you and the global technical community of your competence and reliability
Was designed by global experts. After 12-year lag, updated in late 2017
Has been implemented by over 70,000 organizations globally
Products in global trade do not need to be re-tested or calibrated at import locations if they have been tested by accredited labs.
5
Notes: ISO/IEC 17025:2017 is an ISO standard used by testing and calibration laboratories to show competence in their ability to perform specific tests or calibrations. Accreditation to the standard is a formal recognition of a demonstration of that competence.
ISO/IEC 17025 was initially published in 1999. A revision was added in 2005 and the standard was recently updated in November 2017.
ISO/IEC 17025 enhances the acceptance of products across national borders. By removing the need for additional calibration, testing, medical testing and/or inspection of imports and exports, technical barriers to trade are reduced. In this way, the free-trade goal of a 'product tested or calibrated once and accepted everywhere' can be realized.

## Slide 6: ISO/IEC 17025:2017 Laboratory Process
Notes: This example of the process-based model is a representation of the operational processes of a laboratory, as described in Clause 7 of ISO/IEC 17025:2017.
The most important input to this cycle is customer and other relevant interested parties requirements.
The output of the cycle is customer satisfaction and continual improvement of the management system.

## Slide 7: What is a Lab-MS?
A Laboratory Management System (Lab-MS) determines and continuously improves a laboratory’s quality performance. It lays the foundation for the structure of the laboratory activities.
Below are some elements of a Lab-MS throughout a typical laboratory.
  Document Control – ensures methods are correct and maintains lab records
  Customer Contracts – ensures all test methods and requirements are defined
  Purchasing – selects the correct lab equipment and materials from suppliers
  Personnel – provides training to ensure competency of personnel
  Environment – train people, control incubators, chambers
  Calibration – ensures all measuring devices are accurate and maintained
  Corrective Action – the reporting and analysis of nonconformances
  Risk – identifies potential problems and improvement opportunities
  Audit – reviewing all documentation and processes for compliance
7
Notes: A Lab-MS ensures that an organization understands and improves its quality position.
It provides a framework for implementing improvements and meet regulatory requirements.
A Lab-MS can be applied to any type of laboratory, regardless of size.
Having a accredited Lab-MS means a 3rd party has verified that you have a system for measuring and improving your quality performance.
Some examples of a Lab-MS elements throughout a typical laboratory include document control, customer interactions, personnel training, corrective actions, etc.

## Slide 8: Benefits of ISO/IEC 17025 Accreditation
Internal:
Clearly defined process produces consistent results
Improved operations and communication
Confidence of comparability with other global labs
Process for continual improvement
External:
Global and industry recognition
Build international confidence in your competence, thus business expansion
8
Notes: Improved Operations:
  Remember, the requirements of ISO/IEC 17025 are good laboratory practices.
  Experts from around the world got together and identified the basic elements a laboratory should put in place to have a good management system.
  By using the ISO/IEC 17025 standard to design your management system you are drawing on this expertise.
  You will be designing and installing an effective management system that will lead to improved operations and operating margin.
External Recognition:
  ISO/IEC 17025 is international and is recognized worldwide.
  When you are ISO/IEC 17025 accredited other companies know that you are competent and have a quality management system in place.
  They know you have a corrective action system, a system to handle client complaints and nonconformances.
  They know that you are continually addressing problems and improving your operations and performance.
Market Expansion:
  Some marketplaces require ISO/IEC 17025 accreditation for their laboratories. You will be able to sell services or products to these markets.
  Some marketplaces do not openly require accreditation, but prefer to work with ISO/IEC 17025 accredited providers.
  An unaccredited laboratory faces a definite disadvantage in these markets.

## Slide 9: The Accreditation Process
To become accredited, a laboratory must first implement the requirements of ISO/IEC 17025:2017
Evaluate your current laboratory management system,
Add systems and processes to meet the requirements,
Develop the ‘Documented Information’ for the Lab-MS and your processes,
Implement and use the new management system,
Select an Accreditation Body for the accreditation audit,
Obtain certificate of accreditation,
Celebrate!!
9
Notes: Evaluate your current laboratory management system:
  Many of the requirements of the standard are addressed by practices already in place in the laboratory. These practices may or may not be described in documented information.
  Other requirements of the standard may not be addressed at all and these need to be implemented and documented.
  The standard is designed to bring control and consistency to your laboratory processes. Documenting the processes is part of this control. It helps ensure that people are doing the same thing, to get consistent results.
  You will select an Accreditation Body for the accreditation audit. The auditor from the Accreditation Body will look through your documentation and examine your laboratory activities to ensure that you are meeting the requirements of the standard.
  If you pass that audit, you will be presented with your ISO/IEC 17025 Certificate of Accreditation. Time to Celebrate!

## Slide 10: Requirements
What Does ISO/IEC 17025:2017 Require?
Notes: Now we will review each of the requirements of ISO/IEC 17025:2017…..

## Slide 11: What Are the Requirements?
ISO/IEC 17025:2017 has 8 clauses, but only the last 5 contain requirements.
Clause 1 – Scope
Clause 2 – Normative References
Clause 3 – Terms & Definitions
Clause 4 – General Requirements
Clause 5 – Structural Requirements
Clause 6 - Resource Requirements
Clause 7 – Process Requirements
Clause 8 – Management System Requirements
Notes: The requirements of ISO/IEC 17025:2017 are described in 5 clauses or sections.
  Clause 4 – General Requirements
  Clause 5 – Structural Requirements
  Clause 6 - Resource Requirements
  Clause 7 – Process Requirements
  Clause 8 – Management System Requirements

## Slide 12: 4 General requirements
This clause outlines requirements in 2 sub-clauses:
  4.1 Impartiality
  4.2 Confidentiality
Notes: This clause introduces two sub-clauses relating to certain risks to the organization, 4.1 Impartiality and 4.2 Confidentiality.

## Slide 13: 4.1 Impartiality
There must not be any conflicts of interest:
Commercial
Financial
Relationships
The laboratory must identify risks to impartiality from:
Activities
Relationships
Personnel relationships
Lab must show how it mitigated risks
Management must show a commitment to impartiality.
Notes: The laboratory must ensure that there are no commercial, financial or other pressures that would adversely influence or have the appearance of a conflict of interest by the people involved in the laboratory activities or within relationships with the client.
Regular management communications, discussions in the management review meeting and statements about impartiality in the quality manual or an employee’s handbook are a number of ways to show management commitment to impartiality.
The laboratory will need to identify risks to its impartiality and document how it has identified the risks and how it handles the elimination or minimization of the risks.

## Slide 14: 4.2 Confidentiality
The laboratory and its personnel are responsible for the information obtained or created during the performance of laboratory activities.
All information is considered proprietary information and shall be regarded as confidential, except as required by law.
This means the Laboratory is not to disclose any of this information provided by the clients.
Notes: The laboratory and its personnel are responsible for the information obtained or created during the performance of laboratory activities. All information is considered proprietary information and shall be regarded as confidential, except as required by law.

## Slide 15: 5 Structural requirements
The laboratory must define and document the following about the laboratory operations:
The organizational structure
The management with responsibility
The responsibility of the laboratory personnel.
  Typically, these were best addressed using Operational Charts and Job Descriptions.
The activities of the laboratory
  excluding externally provided activities
All of the above must meet 17025 requirements
Notes: The laboratory must define the organization and management structure of the laboratory and specify the responsibility of the laboratory personnel. Typically, these were best addressed using Operational Charts and Job Descriptions.

## Slide 16: 5 Structural requirements
The laboratory will need to have personnel authorized to ensure that regular communication takes place regarding:
Implementation and improvement of the management system
The effectiveness of the management system
The importance of meeting customers' and regulatory requirements.
Notes: The laboratory will need to have personnel available to manage and improve the management system and ensure that regular communication takes place regarding the effectiveness of the management system and the importance of meeting customers' and regulatory requirements.

## Slide 17: 5 Structural requirements
The laboratory must document its procedures to:
produce consistency of its laboratory activities
ensure that the results are valid.
Typically the laboratory’s processes are documented with:
a Quality Manual
Standard Operating Procedures (SOP’s)
Work Instructions
However the standard does not specifically state the format of the required documentation.
Notes: The laboratory must document its procedures to produce consistency of its laboratory activities and ensure that the results are valid. Typically, a Quality Manual, Standard Operating Procedures (SOP’s) and Work Instructions have been used to document the laboratory’s processes, however the standard does not specifically state the format of the required documentation.

## Slide 18: 6 Resource requirements
The requirements of clause 6 are:
6.1 General
6.2 Personnel
6.3 Facilities and environmental conditions
6.4 Equipment
6.5 Metrological traceability
6.6 Externally provided products and services
Notes: This clause has requirements on the resources required for the laboratory, including personnel, facilities, equipment with traceability and any external products or services required.

## Slide 19: 6.1 General
The laboratory shall have the following:
Staff with the proper training
Adequate facilities to conduct tests
Necessary equipment to assure the validity of results
Support services
Notes: The laboratory shall have the proper staffing of personnel, facilities, necessary equipment that affect the validity of results and the support services needed for the laboratory activities.

## Slide 20: 6.2 Personnel
All personnel of the laboratory, both internal (employees) or external (contractors), shall be competent and work within the structure of the laboratory's management system.
Proper documentation, such as job descriptions, will be required for each job function of the laboratory that details the competence, training, supervision, and authorization of laboratory personnel.
Laboratory management shall communicate the duties, responsibilities and authorities to laboratory personnel through regular meetings or personnel performance reviews.
Notes: All personnel of the laboratory, both internal (employees) or external (contractors), shall be competent and work within the structure of the laboratory's management system.
Proper documentation, such as job descriptions, will be required for each job function of the laboratory that details the competence, training, supervision, and authorization of laboratory personnel.
Laboratory management shall communicate the duties, responsibilities and authorities to laboratory personnel through regular meetings or personnel performance reviews.

## Slide 21: 6.3 Facilities and environmental conditions
The requirements for the facilities and environmental conditions of the laboratory will need to be:
Documented
Controlled
Monitored
Recorded
Any items that adversely affect the validity of results should be minimized or controlled, including:
Dust
Humidity
Electrical supply
Temperature
Vibration
Notes: The requirements for the facilities and environmental conditions of the laboratory will need to be documented, controlled, monitored and recorded and any items that adversely affect the validity of results, such as dust, humidity, electrical supply, temperature, and vibration should be minimized or controlled.

## Slide 22: 6.3 Facilities and environmental conditions
Laboratory activity areas must be:
Defined
Controlled
Separated from areas with incompatible laboratory activities to prevent contamination or interference of the activities.
This includes activities (ex: humidity & temperature) at other (ex: customer) sites or facilities.
Notes: The use of areas affecting laboratory activities must be defined and controlled. There must be separation between areas with incompatible laboratory activities to prevent contamination or interference of the activities.
When the laboratory performs laboratory activities at other sites or facilities, such as a customer site, it shall ensure that the environmental conditions of this document are met. For instance, the temperature and humidity of the customer site will need to be recorded when performing calibrations or testing there.

## Slide 23: 6.4 Equipment
The laboratory shall have access to the proper equipment required for the performance of laboratory activities.
A documented calibration program will need to be established that includes:
Records for all equipment which can influence laboratory activities, including the handling, transport, storage, use and maintenance of equipment
Identification of calibration status, including non-calibrated and out of service
Tamper resistance program to safeguard from unauthorized adjustments that would invalidate the calibration status of the equipment.
Notes: The laboratory shall have access to the proper equipment required for the performance of laboratory activities.
A calibration program will need to be established and records shall be retained for all equipment which can influence laboratory activities. The handling, transport, storage, use and maintenance of equipment will need to be documented.
All equipment will need to be identified to allow the user of the equipment to readily identify the status of calibration. Equipment being used that is not calibrated should be identified as such. Equipment that is suspect shall be taken out of service and isolated.
Void labels or other means to identify tampering will need to be placed on equipment to safeguard from unauthorized adjustments that would invalidate the calibration status of the equipment.

## Slide 24: 6.5 Metrological traceability
The laboratory shall establish and maintain metrological traceability of its measurement results to the International System of Units (SI). **
All calibrations must show traceability through a documented unbroken chain of calibrations, each contributing to the measurement uncertainty for the equipment.
When metrological traceability to the SI units is not technically possible, the laboratory shall demonstrate metrological traceability to an appropriate reference.
**In the US, most laboratories report traceability back to NIST (National Institute of Standards and Technology).
Notes: The laboratory shall establish and maintain metrological traceability of its measurement results and that the measurement results are traceable to the International System of Units (SI).
In the US, most laboratories report traceability back to NIST (National Institute of Standards and Technology). All calibrations must show traceability through a documented unbroken chain of calibrations, each contributing to the measurement uncertainty for the equipment.
When metrological traceability to the SI units is not technically possible, the laboratory shall demonstrate metrological traceability to an appropriate reference.

## Slide 25: 6.6 Externally provided products and services
There should be a documented process for use of external products and services:
They must be suitable for the laboratory activities.
approve and review before and after receipt to ensure the correct items were ordered and received.
Which includes defining the criteria for evaluation
Which includes selection and monitoring of the external providers.
Communication of specific requirements to external providers is typically communicated in the initial purchase order or service agreement.
Notes: The laboratory shall ensure that only suitable externally provided products and services that are used to support the laboratory activities are used.
There should be a documented process for approval and review of the products and services before and after receipt to ensure the correct items were ordered and received. The process should include defining the criteria for evaluation, selection and monitoring of the external providers.
Communication of specific requirements to external providers for the products and services to be provided is typically communicated in the initial purchase order or service agreement.

## Slide 26: 7 Process requirements
The requirements of clause 7 are:
7.1 Review of requests, tenders and contracts
7.2 Selection, verification and validation of methods
7.3 Sampling
7.4 Handling of test or calibration items
7.5 Technical records
7.6 Evaluation of measurement uncertainty
7.7 Ensuring the validity of results
7.8 Reporting of results
7.9 Complaints
7.10 Nonconforming work
7.11 Control of data and information management
Notes: This clause has requirements on the specific processes required for the laboratory, including contract review, method validation, handling of items, measurement uncertainty, reports, complaints, nonconforming work and the control of data and technical records.

## Slide 27: 7.1 Review of requests, tenders and contracts
The laboratory shall have a procedure for the review of requests, tenders and contracts. The procedure shall ensure that:
The requirements are adequately defined and documented
The methods are capable of meeting the customers' requirements
The laboratory has the capability to meet the requirements
Any deviations from the contract are communicated to the customer
Records are kept of all customer correspondence
Notes: The laboratory shall have a procedure for the review of requests, tenders and contracts. The procedure shall ensure that:
The requirements are adequately defined and documented
The methods are capable of meeting the customers' requirements
The laboratory has the capability to meet the requirements
Any deviations from the contract are communicated to the customer
Records are kept of all customer correspondence

## Slide 28: 7.2 Selection, verification and validation of methods
The laboratory shall use appropriate methods and procedures and shall ensure that it uses the latest valid version of a method unless it is not appropriate to do so. The laboratory shall verify that it can properly achieve the required performance the methods before introducing them. Records of the verification shall be retained. Any deviations from the methods must be documented, technically justified, authorized, and accepted by the customer.
Recommended methods used can be:
  Published in national or international standards
  From technical organizations
  Written in relevant scientific texts or journals
  Manuals form the manufacturer of the equipment
  Laboratory-developed methods
Notes: 7.2.1 Selection and verification of methods
The laboratory shall use appropriate methods and procedures and shall ensure that it uses the latest valid version of a method unless it is not appropriate to do so. The laboratory shall verify that it can properly achieve the required performance the methods before introducing them. Records of the verification shall be retained. Any deviations from the methods must be documented, technically justified, authorized, and accepted by the customer.
Recommended methods used can be:
Published in national or international standards
From technical organizations
Written in relevant scientific texts or journals
Manuals form the manufacturer of the equipment
Laboratory-developed methods

## Slide 29: 7.2 Selection, verification and validation of methods
The laboratory shall validate non-standard methods, laboratory-developed methods and standard methods used outside their intended scope. The techniques used for method validation can be any of the following:
  calibration using reference standards or reference materials
  systematic assessment of the factors influencing the result
  testing method robustness through variation of controlled parameters
  comparison of results achieved with other validated methods
  interlaboratory comparisons
If any changes are made to a validated method, the influence of such changes will need to be performed. Records will need to be retained of the validation.
Notes: 7.2.2 Validation of methods
The laboratory shall validate non-standard methods, laboratory-developed methods and standard methods used outside their intended scope. The techniques used for method validation can be any of the following:
calibration using reference standards or reference materials
systematic assessment of the factors influencing the result
testing method robustness through variation of controlled parameters
comparison of results achieved with other validated methods
interlaboratory comparisons
If any changes are made to a validated method, the influence of such changes will need to be performed. Records will need to be retained of the validation.

## Slide 30: 7.3 Sampling
The laboratory shall have a sampling plan and method when it carries out sampling of materials or products for testing or calibration. The sampling method and records shall include:
the selection of samples
the sampling plan
the preparation and treatment of samples
date and time of sampling
identification of the samples
personnel performing the sampling
equipment that was used
environmental conditions
deviations from the sampling plan
Notes: The laboratory shall have a sampling plan and method when it carries out sampling of materials or products for testing or calibration. The sampling method and records shall include:
the selection of samples
the sampling plan
the preparation and treatment of samples
date and time of sampling
identification of the samples
personnel performing the sampling
equipment that was used
environmental conditions
deviations from the sampling plan

## Slide 31: 7.4 Handling of test or calibration items
A procedure will need to be written detailing the handling instructions of test or calibration items. The procedure should list the precautions that shall be taken to avoid contamination or damage to the item during:
  The handling and transporting of the item
  Storage of the item, including any requires environmental conditions
  Preparation for testing or calibration
Any unusual conditions shall be recorded upon receipt of the test or calibration item. When there is doubt about the suitability of an item, the laboratory shall consult with the customer for further instructions before proceeding and shall record the results of this consultation.
The laboratory shall have a system for the unique identification of the items which is to be retained while the items are under the responsibility of the laboratory.
Notes: A procedure will need to be written detailing the handling instructions of test or calibration items. The procedure should list the precautions that shall be taken to avoid contamination or damage to the item during:
The handling and transporting of the item
Storage of the item, including any requires environmental conditions
Preparation for testing or calibration
Any unusual conditions shall be recorded upon receipt of the test or calibration item. When there is doubt about the suitability of an item, the laboratory shall consult with the customer for further instructions before proceeding and shall record the results of this consultation.
The laboratory shall have a system for the unique identification of the items which is to be retained while the items are under the responsibility of the laboratory.

## Slide 32: 7.5 Technical records
The laboratory shall ensure that technical records for each laboratory activity contain sufficient information to enable the activity to be repeated under conditions as close as possible to the original.
  The technical records shall include the date and the identity of personnel responsible for each laboratory activity and for checking data and results.
Original observations, data and calculations shall be recorded at the time they are made and must have identifying information, (job number, serial number, etc.) to tie the information to the specific item.
The laboratory shall ensure that amendments or changes to technical records shall be retained, including:
  the date of the change
  what was changed
  the personnel responsible for the change.
Notes: The laboratory shall ensure that technical records for each laboratory activity contain sufficient information to enable the activity to be repeated under conditions as close as possible to the original.
  The technical records shall include the date and the identity of personnel responsible for each laboratory activity and for checking data and results.
Original observations, data and calculations shall be recorded at the time they are made and must have identifying information, (job number, serial number, etc.) to tie the information to the specific item.
The laboratory shall ensure that amendments or changes to technical records shall be retained, including:
  the date of the change
  what was changed
  the personnel responsible for the change.

## Slide 33: 7.6 Evaluation of measurement uncertainty
Laboratories shall identify the significant contributions to measurement uncertainty for all items listed on its Scope of Accreditation, even if the measurement uncertainty is not reported to the customer. Those records will need to be retained.
If a well-recognized test method is used, such as an ASTM test method, the laboratory is considered to have satisfied this clause by following the test method and reporting instructions.
Notes: Laboratories shall identify the significant contributions to measurement uncertainty for all items listed on its Scope of Accreditation, even if the measurement uncertainty is not reported to the customer. Those records will need to be retained.
If a well-recognized test method is used, such as an ASTM test method, the laboratory is considered to have satisfied this clause by following the test method and reporting instructions.

## Slide 34: 7.7 Ensuring the validity of results
The laboratory will need to document how they monitor the validity of results. Data from the monitoring activities will need to be analyzed and used to improve the laboratory's activities. The analysis results and the actions taken will need to be documented.
This monitoring shall include any of the following items:
  use of reference materials or quality control materials
  use of alternative instrumentation that has been calibrated to provide traceable results
  functional check(s) of measuring and testing equipment
  use of check or working standards with control charts, where applicable
  intermediate checks on measuring equipment
  replicate tests or calibrations using the same or different methods
  retesting or recalibration of retained items
  correlation of results for different characteristics of an item
  review of reported results
  intra-laboratory comparisons
Notes:   The laboratory will need to document how they monitor the validity of results. Data from the monitoring activities will need to be analyzed and used to improve the laboratory's activities. The analysis results and the actions taken will need to be documented.
  use of reference materials or quality control materials;
  use of alternative instrumentation that has been calibrated to provide traceable results;
  functional check(s) of measuring and testing equipment;
  use of check or working standards with control charts, where applicable;
  intermediate checks on measuring equipment;
  replicate tests or calibrations using the same or different methods;
  retesting or recalibration of retained items;
  correlation of results for different characteristics of an item;
  review of reported results;
  intralaboratory comparisons;

## Slide 35: 7.8 Reporting of results
There will need to be an inspection process to review and authorize the results prior to release. The results shall be provided accurately and clearly in a test report or a calibration certificate. Each report shall include the following information:
a title (e.g. “Test Report”, “Calibration Certificate” or “Report of Sampling”);
the name and address of the laboratory and the location of performance of the laboratory activities
unique identification that all its components are recognized as a portion of a complete report and a clear identification of the end
the name and contact information of the customer
identification of the method used
unambiguous identification and the condition of the item
the date of performance of the laboratory activity and the date of issue of the report
the results with, where appropriate, the units of measurement
identification of the person authorizing the report
any information agreed with the customer
a statement that the results relate only to the items tested, calibrated or sampled
a statement specifying that the report shall not be reproduced except in full without approval of the laboratory.
Amendments to a report after issue shall be made only in the form of an additional document, or data transfer, which includes the statement “Amendment to Report, serial number... [or as otherwise identified]”, or an equivalent form of wording.
Notes: There will need to be an inspection process to review and authorize the results prior to release. The results shall be provided accurately and clearly in a test report or a calibration certificate. Each report shall include the following information:
a title (e.g. “Test Report”, “Calibration Certificate” or “Report of Sampling”);
the name and address of the laboratory and the location of performance of the laboratory activities
unique identification that all its components are recognized as a portion of a complete report and a clear identification of the end
the name and contact information of the customer
identification of the method used
unambiguous identification and the condition of the item
the date of performance of the laboratory activity and the date of issue of the report
the results with, where appropriate, the units of measurement
identification of the person authorizing the report
any information agreed with the customer
a statement that the results relate only to the items tested, calibrated or sampled
a statement specifying that the report shall not be reproduced except in full without approval of the laboratory.
Amendments to a report after issue shall be made only in the form of an additional document, or data transfer, which includes the statement “Amendment to Report, serial number... [or as otherwise identified]”, or an equivalent form of wording.

## Slide 36: 7.9 Complaints
The laboratory will need to have a documented process to receive, evaluate and make decisions on complaints, both internal and external.
Upon receipt of a complaint, the laboratory shall investigate whether the complaint relates to laboratory activities that it is responsible for and, if so, shall deal with it.
  (Not all complaints are credible. The laboratory must decide which complaints to act upon.)
The process for handling complaints must include:
  a description of the process for receiving, validating, investigating the complaint, and deciding what actions are to be taken in response to it,
  tracking and recording complaints,
  including actions undertaken to resolve them,
  and ensuring that any appropriate action is taken.
  documented records must be kept of the decisions.
Whenever possible, the laboratory shall acknowledge receipt of the complaint, and provide the complainant with progress reports and the outcome. (New requirement for ISO/IEC 17025:2017)
Notes: The laboratory will need to have a documented process to receive, evaluate and make decisions on complaints, both internal and external.
Upon receipt of a complaint, the laboratory shall investigate whether the complaint relates to laboratory activities that it is responsible for and, if so, shall deal with it.
  (Not all complaints are credible. The laboratory must decide which complaints to act upon.)
The process for handling complaints must include:
  a description of the process for receiving, validating, investigating the complaint, and deciding what actions are to be taken in response to it,
  tracking and recording complaints,
  including actions undertaken to resolve them,
  and ensuring that any appropriate action is taken.
  documented records must be kept of the decisions.
Whenever possible, the laboratory shall acknowledge receipt of the complaint, and provide the complainant with progress reports and the outcome. (New requirement for ISO/IEC 17025:2017)

## Slide 37: 7.10 Nonconforming work
A documented procedure is required that details the actions taken when any aspect of the laboratory activities or results do not conform to its own procedures or the agreed requirements of the customer. Detailed records must be kept of all occurrences and decisions made in the nonconformances.
The procedure shall ensure that:
  the responsibilities and authorities for the management of the nonconforming work are defined,
  any actions are based upon the risk levels established by the laboratory,
  an evaluation is made of the significance of the nonconforming work, including an impact analysis on previous results,
  a decision is taken on the acceptability of the nonconforming work,
  where necessary, the customer is notified and work is recalled,
  the responsibility for authorizing the resumption of work is defined.
Notes: A documented procedure is required that details the actions taken when any aspect of the laboratory activities or results do not conform to its own procedures or the agreed requirements of the customer. Detailed records must be kept of all occurrences and decisions made in the nonconformances.
The procedure shall ensure that:
  the responsibilities and authorities for the management of the nonconforming work are defined,
  any actions are based upon the risk levels established by the laboratory,
  an evaluation is made of the significance of the nonconforming work, including an impact analysis on previous results,
  a decision is taken on the acceptability of the nonconforming work,
  where necessary, the customer is notified and work is recalled,
  the responsibility for authorizing the resumption of work is defined.

## Slide 38: 7.11 Control of data and information management
The laboratory information management system used for the collection, processing, recording, reporting, storage or retrieval of data shall be validated for functionality by the laboratory before introduction. Any changes or modifications must be validated before implementation.
  Commercial off-the-shelf software in general use within its designed application range, such as Excel, Access, etc. can be considered to be sufficiently validated.
  (For example, in an Excel formula, you do not need to validate that the addition of the values in cell A1 and A2 is correct. That is the intended function of Excel. You do need to validate that the formula is accessing the correct cells in the formula. If the wrong cells are put into the formula, you will get the wrong result.)
The laboratory information management system must
  be protected from unauthorized access,
  be safeguarded against tampering,
  be operated in an environment that safeguards the accuracy of manual recording,
  ensure the integrity of the data.
Notes: The laboratory information management system used for the collection, processing, recording, reporting, storage or retrieval of data shall be validated for functionality by the laboratory before introduction. Any changes or modifications must be validated before implementation.
  Commercial off-the-shelf software in general use within its designed application range, such as Excel, Access, etc. can be considered to be sufficiently validated.
  For example, in an Excel formula, you do not need to validate that the addition of the values in cell A1 and A2 is correct. That is the intended function of Excel. You do need to validate that the formula is accessing the correct cells in the formula. If the wrong cells are put into the formula, you will get the wrong result.
The laboratory information management system must
  be protected from unauthorized access,
  be safeguarded against tampering,
  be operated in an environment that safeguards the accuracy of manual recording,
  ensure the integrity of the data.

## Slide 39: 8 Management system requirements
The requirements for clause 8 are:
8.1 Options
8.1.1 General
8.1.2 Option A
8.1.3 Option B
8.2 Management system documentation (Option A)
8.3 Control of management system documents (Option A)
8.4 Control of records (Option A)
8.5 Actions to address risks and opportunities (Option A)
8.6 Improvement (Option A)
8.7 Corrective actions (Option A)
8.8 Internal audits (Option A)
8.9 Management reviews (Option A)
Notes: This clause has requirements on the specific management system requirements for the laboratory, including documentation and record management, addressing risks and improvement opportunities, corrective actions, internal audits and management reviews.

## Slide 40: 8.1 Options
The laboratory will need to establish, document, implement and maintain a management system that is capable of supporting and demonstrating the consistent achievement of the requirements of ISO/IEC 17025 and assuring the quality of the laboratory results. In addition to meeting the requirements of Clauses 4 to 7, the laboratory shall implement a management system in accordance with Option A or Option B.
Option A lists the minimum requirements for implementation of a management system in a laboratory. All those requirements of ISO 9001 that are relevant to the scope of laboratory activities that are covered by the management system. Laboratories that comply with Clauses 4 to 7 and implement Option A of Clause 8 will therefore also operate generally in accordance with the principles of ISO 9001.
Option B is for those laboratories that are ISO 9001 certified and wish to achieve the ISO/IEC 17025 accreditation. It states that the management system of ISO 9001 will fulfill the requirements of Clause 8 of the ISO/IEC 17025 standard.
  Conformity to the ISO 9001 management system does not demonstrate the competence of the laboratory to produce technically valid data and results. This is accomplished through the consistent fulfilment of the requirements of Clauses 4 to 7.
Both options are intended to achieve the same result in the performance of the management system and compliance with Clauses 4 to 7.
Notes: The laboratory will need to establish, document, implement and maintain a management system that is capable of supporting and demonstrating the consistent achievement of the requirements of ISO/IEC 17025 and assuring the quality of the laboratory results. In addition to meeting the requirements of Clauses 4 to 7, the laboratory shall implement a management system in accordance with Option A or Option B.
Option A lists the minimum requirements for implementation of a management system in a laboratory. All those requirements of ISO 9001 that are relevant to the scope of laboratory activities that are covered by the management system. Laboratories that comply with Clauses 4 to 7 and implement Option A of Clause 8 will therefore also operate generally in accordance with the principles of ISO 9001.
Option B is for those laboratories that are ISO 9001 certified and wish to achieve the ISO/IEC 17025 accreditation. It states that the management system of ISO 9001 will fulfill the requirements of Clause 8 of the ISO/IEC 17025 standard.
  Conformity to the ISO 9001 management system does not demonstrate the competence of the laboratory to produce technically valid data and results. This is accomplished through the consistent fulfilment of the requirements of Clauses 4 to 7.
Both options are intended to achieve the same result in the performance of the management system and compliance with Clauses 4 to 7.

## Slide 41: 8.2 Management system documentation (Option A)
Laboratory management shall establish, document, and maintain policies and objectives for the fulfilment of the purposes of ISO/IEC 17025 and shall ensure that the policies and objectives are acknowledged and implemented at all levels of the laboratory organization.
  The policies and objectives of the laboratory shall address the competence, impartiality and consistent operation of the laboratory.
  Laboratory management shall be commitment to the development and implementation of the management system and will need to show evidence of improving its effectiveness.
  All documentation, processes, and records shall be included in, referenced from, or linked to the management system.
Notes: Laboratory management shall establish, document, and maintain policies and objectives for the fulfilment of the purposes of ISO/IEC 17025 and shall ensure that the policies and objectives are acknowledged and implemented at all levels of the laboratory organization.
  The policies and objectives of the laboratory shall address the competence, impartiality and consistent operation of the laboratory.
  Laboratory management shall be commitment to the development and implementation of the management system and will need to show evidence of improving its effectiveness.
  All documentation, processes, and records shall be included in, referenced from, or linked to the management system.

## Slide 42: 8.3 Control of management system documents (Option A)
The Lab-MS includes the documented information required by the ISO/IEC 17025:2017 standard and the documented information determined to be necessary for an effective Lab-MS. Documented information must be controlled to ensure that it is available and suitable for use, where and when it is needed and it is adequately protected.
  Many companies will use a “Master List” to list the current revision and location of each document.
  Recording the distribution of documents is important; if a document is revised all previous revisions of the document must be replaced. This is only possible if you know where all those copies are.
  Documented information from external sources are controlled by the owner of the documents. The external documents must be regularly reviewed to ensure that the latest revision is being used.
The laboratory will need to ensure that:
  documents are approved for adequacy prior to issue by authorized laboratory personnel;
  documents are periodically reviewed, typically annually, and updated as necessary;
  changes and the current revision status of documents are identified. A revision history page works well to document this.
  relevant versions of applicable documents are available at points of use and, where necessary, their distribution is controlled;
  documents have a unique identification;
  the unintended use of obsolete documents is prevented, and suitable identification is applied to them if they are retained for any purpose.
Notes: The Lab-MS includes the documented information required by the ISO/IEC 17025:2017 standard and the documented information determined to be necessary for an effective Lab-MS. Documented information must be controlled to ensure that it is available and suitable for use, where and when it is needed and it is adequately protected.
Many companies will use a “Master List” to list the current revision and location of each document.
Recording the distribution of documents is important; if a document is revised all previous revisions of the document must be replaced. This is only possible if you know where all those copies are.
Documented information from external sources are controlled by the owner of the documents. The external documents must be regularly reviewed to ensure that the latest revision is being used.
The laboratory will need to ensure that:
  documents are approved for adequacy prior to issue by authorized laboratory personnel;
  documents are periodically reviewed, typically annually, and updated as necessary;
  changes and the current revision status of documents are identified. A revision history page works well to document this.
  relevant versions of applicable documents are available at points of use and, where necessary, their distribution is controlled;
  documents have a unique identification;
  the unintended use of obsolete documents is prevented, and suitable identification is applied to them if they are retained for any purpose.

## Slide 43: 8.4 Control of records (Option A)
The laboratory must keep all records related to the laboratories operations to demonstrate fulfilment of the requirements in this document. The laboratory will need to ensure that all of its records are stored, backed-up and protected.
  A table or spreadsheet works well for listing all your quality records, where they are generated and stored, how long they are stored and who is responsible for them.
The records will need to be readily retrievable and information on how long the records are kept will need to be documented.
  ISO/IEC 17025 does not specify how long you have to keep your records, but remember that the auditor will want to see six months to a year of records for an initial assessment and at least two years worth of records between scheduled assessments.
Notes: The laboratory must keep all records related to the laboratories operations to demonstrate fulfilment of the requirements in this document. The laboratory will need to ensure that all of its records are stored, backed-up and protected.
  A table works well for listing all your quality records, where they are generated and stored, how long they are stored and who is responsible for them.
The records will need to be readily retrievable and information on how long the records are kept will need to be documented.
  ISO/IEC 17025 does not specify how long you have to keep your records, but remember that the auditor will want to see six months to a year of records for an initial assessment and at least two years worth of records between scheduled assessments.

## Slide 44: 8.5 Actions to address risks and opportunities (Option A)
Actions to address the risks and opportunities are determined in order to integrate and implement them into the Lab-MS processes and to evaluate the effectiveness of these actions.
The risks and opportunities addressed by the laboratory in order to ensure that the Lab-MS can:
  achieve its intended results,
  enhance opportunities to achieve the objectives of the laboratory;
  prevent, or reduce, undesired effects, and
  achieve continual improvement.
Not all risks or opportunities that are identified need to be addressed. The Laboratory must evaluate each risk or opportunity and determine the overall impact it will have on the laboratories activities.
A formal risk management process is not required for the fulfillment of the ISO/IEC 17025 standard. However, an auditor will want to see risk assessments through:
  continuous improvement activities,
  management communications,
  corrective actions,
  other documentation that shows implementation and evaluation of those activities.
Notes: Actions to address the risks and opportunities are determined in order to integrate and implement them into the Lab-MS processes and to evaluate the effectiveness of these actions.
The risks and opportunities addressed by the laboratory in order to ensure that the Lab-MS can:
  achieve its intended results,
  enhance opportunities to achieve the objectives of the laboratory;
  prevent, or reduce, undesired effects, and
  achieve continual improvement.
Not all risks or opportunities that are identified need to be addressed. The Laboratory must evaluate each risk or opportunity and determine the overall impact it will have on the laboratories activities.
A formal risk management process is not required for the fulfillment of the ISO/IEC 17025 standard. However, an auditor will want to see risk assessments through:
  continuous improvement activities,
  management communications,
  corrective actions,
  other documentation that shows implementation and evaluation of those activities.

## Slide 45: 8.6 Improvement (Option A)
Your laboratory must determine and select opportunities for improvement and implement those actions. Opportunities for improvement can be identified through a variety of methods including:
  the review of the operational procedures and policies,
  overall objectives,
  audit results,
  corrective actions,
  management review,
  risk assessment,
  proficiency testing results.
Your laboratory must seek feedback from your customers, both positive and negative, and the feedback must be analyzed to improve the management system, laboratory activities and customer service. The methods for obtaining this information can include:
  customer satisfaction or opinion surveys,
  communication records and
  review of reports with customers.
Notes: Your laboratory must determine and select opportunities for improvement and implement those actions. Opportunities for improvement can be identified through a variety of methods including:
  the review of the operational procedures,
  the use of the policies,
  overall objectives,
  audit results,
  corrective actions,
  management review,
  risk assessment,
  proficiency testing results.
Your laboratory must seek feedback from your customers, both positive and negative, and the feedback must be analyzed to improve the management system, laboratory activities and customer service. The methods for obtaining this information can include:
  customer satisfaction or opinion surveys,
  communication records and
  review of reports with customers.

## Slide 46: 8.7 Corrective actions (Option A)
Corrective actions shall be appropriate to the effects of the nonconformities encountered. Nonconformities that carry more risk should have a more thorough corrective action applied.
When a nonconformity occurs, the laboratory shall:
  take action to control and correct it;
  address the consequences;
  review and analyze the nonconformity;
  determine the causes of the nonconformity;
  determine if similar nonconformities exist, or could potentially occur;
  make changes to the management system, if necessary.
Documentation shall be retained as evidence of:
  the nature of the nonconformities, cause(s) and any subsequent actions taken;
  the results of any corrective action.
Notes: Corrective actions shall be appropriate to the effects of the nonconformities encountered. Nonconformities that carry more risk should have a more thorough corrective action applied.
When a nonconformity occurs, the laboratory shall:
  take action to control and correct it;
  address the consequences;
  review and analyze the nonconformity;
  determine the causes of the nonconformity;
  determine if similar nonconformities exist, or could potentially occur;
  make changes to the management system, if necessary.
Documentation shall be retained as evidence of:
  the nature of the nonconformities, cause(s) and any subsequent actions taken;
  the results of any corrective action.

## Slide 47: 8.8 Internal audits (Option A)
Your laboratory must conduct internal audits to ensure that the Lab-MS conforms to the laboratories own requirements, the requirements of the ISO/IEC 17025 standard and is effectively implemented and maintained.
Your laboratory must:
  plan, establish, implement and maintain an audit program
  define the audit criteria and scope for each audit,
  ensure that the results of the audits are reported to relevant management,
  take necessary correction and corrective actions without undue delay,
  retain documented information as evidence of the implementation of the audit program and the audit results.
Notes: Your laboratory must conduct internal audits to ensure that the Lab-MS conforms to the laboratories own requirements, the requirements of the ISO/IEC 17025 standard and is effectively implemented and maintained.
Your laboratory must:
  plan, establish, implement and maintain an audit program
  define the audit criteria and scope for each audit,
  ensure that the results of the audits are reported to relevant management,
  take necessary correction and corrective actions without undue delay,
  retain documented information as evidence of the implementation of the audit program and the audit results

## Slide 48: 8.9 Management reviews (Option A)
Top management in your laboratory must review the Lab-MS, at planned intervals, to ensure that the Lab-MS continues to be suitable, adequate and effective.
Management reviews must be conducted with consideration given to the following inputs:
  Changes in external and internal issues that are relevant to the laboratory,
  Fulfilment of the quality objectives,
  The suitability of the existing policies and procedures,
  The status of actions from previous management reviews,
  Audit results,
  Nonconformities and corrective actions,
  Any assessments by external bodies
  Management reviews must be planned and conducted with consideration given to the following inputs:
  Changes in the volume or type of work for the laboratory,
  Customer satisfaction,
  Complaints, both internal and external,
  The effectiveness of any implemented improvements,
  Adequacy of resources required for the laboratory,
  Actions taken to address risks and opportunities,
  Outcomes of the assurance of the validity of results, including proficiency tests,
  And any other relevant factors, such as monitoring results and training.
The outputs of the management review must include decisions and actions related to:
  the effectiveness of the management system;
  improvement of the laboratory activities related to the fulfillment of the ISO/IEC 17025 standard;
  resources need for the improvement of the management system and laboratory activities;
  any need for changes to the management system.
Notes: Top management in your laboratory must review the Lab-MS, at planned intervals, to ensure that the Lab-MS continues to be suitable, adequate and effective.
Management reviews must be conducted with consideration given to the following inputs:
  Changes in external and internal issues that are relevant to the laboratory,
  Fulfilment of the quality objectives,
  The suitability of the existing policies and procedures,
  The status of actions from previous management reviews,
  Audit results,
  Nonconformities and corrective actions,
  Any assessments by external bodies
  Management reviews must be planned and conducted with consideration given to the following inputs:
  Changes in the volume or type of work for the laboratory,
  Customer satisfaction,
  Complaints, both internal and external,
  The effectiveness of any implemented improvements,
  Adequacy of resources required for the laboratory,
  Actions taken to address risks and opportunities,
  Outcomes of the assurance of the validity of results, including proficiency tests,
  And any other relevant factors, such as monitoring results and training.
The outputs of the management review must include decisions and actions related to:
  the effectiveness of the management system;
  improvement of the laboratory activities related to the fulfillment of the ISO/IEC 17025 standard;
  resources need for the improvement of the management system and laboratory activities;
  any need for changes to the management system.

## Slide 49: The ISO/IEC 17025:2017 Requirements
With this presentation, we have covered the requirements of ISO/IEC 17025:2017.
The 5 clauses of the international standard are:
Clause 4 – General Requirements
Clause 5 – Structural Requirements
Clause 6 - Resource Requirements
Clause 7 – Process Requirements
Clause 8 – Management System Requirements
Notes: With this presentation, we have covered the requirements of ISO/IEC 17025:2017.
  Clause 4 – General Requirements
  Clause 5 – Structural Requirements
  Clause 6 - Resource Requirements
  Clause 7 – Process Requirements
  Clause 8 – Management System Requirements

## Slide 50: Next Steps
Determine timeline for implementation
Perform Gap Analysis Review how your existing quality system fits into the ISO/IEC 17025:2017 format
Put together an implementation plan and timeline
Identify an Accreditation Body
Implement the system
The information you get from your Gap analysis, combined with a list of key laboratory activities should give you a good idea of what you will need to do to implement the standard. Then you can build an outline for the tasks you need to do and the documents you will need to get into place for your accreditation audit.
Notes: Each organization will have its own way to approach the implementation
Performing a gap analysis of your current system versus the requirements of the Standard will give you a task list to work from to plan your implementation project.
The information you get from your Gap analysis, combined with a list of key laboratory activities should give you a good idea of what you will need to do to implement the standard. Then you can build an outline for the tasks you need to do and the documents you will need to get into place for your accreditation audit.

## Slide 51: Next Steps
Run the system for 3 months to gather records, perform internal audits, have management review and make improvements
Have your Accreditation Audit
Correct any non-conformances
Obtain your certificate of Accreditation
Celebrate your Accreditation!
Maintain your system with regular internal audits, management reviews and surveillance audits from your Accreditation Body
Be successful!
Notes: Run the system for several months to gather records, perform internal audits, have the management review and make improvements.
Have your Accreditation Audit and correct any non-conformances found in the audit.
After your initial accreditation, you will need to maintain your Lab-MS with regular internal audits, management reviews and surveillance audits from your Accreditation Body.
Be successful!